Document Authorization:

|  | Name | Date | Signature |
| --- | --- | --- | --- |
| Owner | Emily Rydin | 15Oct18 | See Signature on File |
| Operational Management | Pierre Barratt | 16Oct18 | See Signature on File |
| Quality Assurance | Stacey Kaufusi | 19Oct18 | See Signature on File |

Changes Since Last Version:

| Section | Summary of Changes | Change Control Number |
| --- | --- | --- |
| All | Updated header, footer and formatting. | PCC18007 |
| 3 | Internal references updated to reflect content of SOP, obsolete PFRM013 (208956) Balance Logbook, and include PFRM021 (208868) General Equipment Use Log and 209042 Balance Drift Log | PCC18014 |
| 5 | Quality Assurance responsibilities removed as not applicable to this SOP. |  |

- PURPOSE

The purpose of this document is to describe the calibration, use and maintenance of analytical balances.

- SCOPE

This procedure applies to calibration, use and maintenance of all laboratory balances for basic weighing operations, within the Therapeutic GMP Manufacturing Suite and Analytical Testing Laboratory in Petaluma, CA.

- INTERNAL REFERENCES

| Document ID | Title |
| --- | --- |
| PPOL001 (208849) | Quality Policy |
| PSOP012 (208815) | Supplier Qualification Program |
| PSOP130 (208943) | Commissioning of Manufacturing and Test Equipment |
| PFRM021 (208868) | General Equipment Use Log |
| FM-209042 | Balance Drift Log |

- EXTERNAL REFERENCES

| Document ID | Title |
| --- | --- |
| Mettler Toledo AX and MX/UMX Balances User’s Manual | Operating Instructions Mettler Toledo AX and MX/UMX Balances (© 2004) |
| Mettler Toledo Excellence Plus Precision Balances XP Models User’s Manual | Operating Instructions Mettler Toledo Excellence Plus Precision Balances XP Models- Part 1 (© 2012) |
| Mettler Toledo RS-P42 Printer Manual | Operating Instructions Mettler Toledo RS-P42 Printer (© 2009) |
| Mettler Toledo LC-P45 Printer Manual | Operating Instructions Mettler Toledo LC-P45 Printer (© 2009) |
| ASTM E617-13 | Standard Specification for Laboratory Weights and Precision Mass Standards |
| USP <41> | Weights and Balances |
| USP <1251> | Weighing on an Analytical Balance |
| NISTIR 6919 | Recommended Guide for Determining and Reporting Uncertainties for Balances and Scales. |
| ICH Q7 (API) | Good Manufacturing Practice Guidance for Active Pharmaceutical Ingredients |

- RESPONSIBILITIES

| Job Function and/or Department | Responsibility |
| --- | --- |
| Operational Employees | - Ensure that an average weight (for use when performing drift checks) is generated on all balances following the annual balance calibration. The average weight is to be calculated on not less than 10 replicate weighing’s of an external weight. Printouts are to be attached to the Balance Drift Log (209042) and checked accordingly. - Ensure that an internal calibration and performance drift check has been performed prior to use. Should the tests have not been performed, then perform such tests and ensure that they are documented on the Balance Drift Log (209042). - Ensure that the time and date stamps on the printouts are accurate and correlate with one another. - Time and date stamps on printouts should refer to the time in the current time zone excluding changes due to daylight savings (i.e. UTC – 8 Pacific Standard Time). - Ensure internal calibration and performance drift check is to be performed prior to use each day and a “GLP compliant” *printout is to be included on the balance printout retained within the Equipment Logbook. - Ensure that the balance is within its calibration date. - Notify operational management if the performance drift check fails to meet the required specifications. Investigate any failures in performance of the balance in performance drift check weighing and follow appropriate actions. |
| Instrument Owner/ Supervisor | - Ensure that the annual calibration is performed on each balance. - Ensure that balance weights used for the performance of drift checking are re-calibrated on a two-year basis. - Review the balance drift logs and results monthly during logbook review to assess any patterns in balance drift. - Perform a periodic review of the balance logbooks on a quarterly basis. Investigate any failures in performance of the balance in performance drift check weighing and follow appropriate actions. |

- DEFINITION

| Term | Definition |
| --- | --- |
| Certified Check Weight | A class 1 weight is used for the performance check of the balance. No adjustments to the balance are made with this weight. |
| Applied Mass | The certified value assigned to the check weight, typically assigned by the metrology department or external accredited calibration facility. NOTE: Applied mass is different from the nominal mass of the weight. (E.g., for a weight with a nominal mass of 10 grams, the applied mass might be 9.99987 grams.) |
| ASTM Class 1 Balance weight | Used as a reference standard in calibrating other weights and is appropriate for calibrating high-precision analytical balances with a readability of as low as 0.1 mg to 0.01 mg |
| Internal Adjustment or Calibration | The process of using an internal balance weight, to adjust or set the span of the balance. |
| Minimum Weight (Mmin) | The lowest weight that meets USP precision guidelines, for being “accurately weighed”. The minimum weight calculation is defined in the USP as the following: Mmin = 2000 x s Where s = standard deviation of not fewer than 10 replicates of a test weight. |
| Performance Drift Check | The process of using a weight, external to the balance to verify accuracy of the internal or external calibration, and to verify drift of the balance. |

- PROCEDURE

  - Calibration of Balances

All analytical balances are subjected to an annual planned maintenance service visit (and additionally if their location is changed) by an accredited external organization who are approved as per PSOP012 (208815) Supplier Qualification Program. During this calibration the measurement uncertainty, linearity, repeatability and off-center/eccentricity loading will be determined as shown below.

    - Measurement Uncertainty

Using ASTM Class 1 balance weight of appropriate mass, Perform 10 replicate weighing’s. Calculate standard deviation(S) of the 10-replicate weighing’s, using the following formula:

FIGURE 1:  STANDARD DEVIATION CALCULATION

Where:

- s = standard deviation

- x = each value in the sample

- x = the mean of the values

- N= the number of values (10)

Measurement uncertainty (MU) is calculated as being two times the standard deviation of ten replicate weighing’s divided by the mean of the ten readings for that weight:

FIGURE 2: MEASUREMENT UNCERTAINTY CALCULATION

MU = (2 × s)Mean Weight

Hence, the minimum weight (Mmin) can be calculated as follows:

FIGURE 3:  MINIMUM WEIGHT CALCULATION

Mmin =MU × Mean Weight × 1000

    - Linearity

This test involves at least ten weighing’s through and to the maximum of the range under test. The applied load, indicated value, and maximum difference between the two are to be recorded on the certificate.

    - Repeatability

Repeatability is assessed by weighing one test weight not less than 10 times. Repeatability is satisfactory if two times the standard deviation of the weighed value, divided by the desired smallest net weight, does not exceed 0.10%. If the repeatability obtained is smaller than 0.41d (where d is the scale interval, replace this standard deviation with 0.41d).

    - Off-Center/Eccentric Loading

To be performed by placing a load between ¼ and 1/3 of the range under test on the balance pan. The load is then moved to the various positions as shown below and the maximum deviation from the center recorded.

A

B

C

E

D

FIGURE 4:  OFF- CENTER / ECCENTRIC LOADING

  - Use of Balances

    - General Requirements

      - The balance should not be turned off at the mains except when necessary. When not in use it may be turned off at the control panel.

      - If switched off at the mains, allow at least 12 hours equilibration before re-use or as stated by the manufacturer.

      - The balance pan or platform must be clean, free of material and level prior to use. The level of the balance is checked by using the spirit level and the bubble should be within the marked circle on the level. The balance should not rock when pressed on any corner, as this will demonstrate that one or more feet are not in contact with the work surface.

      - For balances with enclosure doors, close the doors prior to performing the internal calibration, taring or recording measurements.

      - Do not place magnetic items (such as stir bars) on the balance.

      - Clean the balance and surrounding area before leaving the work area or leave notification on the balance that a weighing operation is in progress.

      - Care should be taken to ensure the balance is not knocked, and that items are not dropped on the balance pan.

      - Prior to any weighing taking place on the day of use the internal calibration and balance drift (as per 7.2.2) will be performed. Should the instrument perform a further automated internal calibration (due to time elapsed, temperature, and/or humidity changes) during your analysis this calibration printout should be retained on the printout and include in the experimental write up.

      - Ensure that the balance selected is appropriate for the weighing task to be performed. The minimum weight for each balance is calculated by the accredited external organization during the annual calibration visit and is displayed on each balance. Wherever possible the minimum weight of sample (not tare or gross weight) must be equal to or larger than the calculated minimum weight, to meet the weighing tolerance requirements of USP <41>. If it is not possible to weigh above the minimum balance weight, then justification should be given as to the acceptability of the possible variation and its impact on the testing.

      - Ideally, standard or sample weighing’s of less than 10 mg should be avoided. Weights of at least 25mg if possible, should be used and measured using an appropriate balance according to its minimum balance weight.

      - Where possible samples and chemical standards should be weighed in clean, dry, empty containers with a neck wide enough to allow easy access with a spatula, e.g., Glass weighing boat, small beaker.

      - Weigh liquids into small narrow necked or closed container to minimize evaporation.

      - Care should be taken if using plastic containers due to possible effects of static electricity.

      - Where possible weights should be printed out using an attached printer or electronically downloaded to an electronic lab notebook system.

    - Daily Checks

      - In addition to performing an internal balance calibration, all balances are checked for measurement of balance drift on the day of use, prior to any weighing taking place, using the following procedure:

NOTE: It is critical that mass standards are at thermal equilibrium with the balance and surrounding environment prior to check weighing/calibration.

      - Ensuring the balance is clean and level, and the doors are closed perform the internal calibration.

      - Tare the balance. Either with plastic forceps or cotton gloves carefully place the check weight in the center of the pan. Allow the reading to settle and then print the weight. Retain the printout for inclusion in the balance logbook. Manually record the value in the balance logbook and check its acceptability according to the tolerance values given in the logbook and in the table below (Tolerance values are 0.05% of the average mass of the check weight and are based on the USP acceptance criteria for balance sensitivity):

TABLE 1:  USP ACCEPTANCE CRITERIA FOR BALANCE SENSITIVITY

| Balance Type | Suggested Mass of Check Weight | Allowable Tolerance for Daily Drift Check (mean ± g) |
| --- | --- | --- |
| 2 place top pan | 200 g | 0.10 |
| 4 place analytical | 50 g | 0.0250 |
| 5 place analytical | 50 g | 0.02500 |
| 6 place analytical microbalance | 1 g | 0.000500 |

      - If the reading does not meet the tolerance value required, remove the check weight. Repeat the internal calibration carefully clean the weight with a lint-free cloth moistened with a small amount of solvent, allow to dry, and repeat the weighing.

      - Upon failure to meet the required tolerance value a second time, the balance should not be used, a “DO NOT USE” label attached, and further investigation be performed.

      - Document the drift check within the Balance Drift Log.

      - The acceptability is assessed against the tolerance supplied on the appropriate balance control chart.

      - Immediately following annual calibration, a balance weight of appropriate mass* is to be weighed ten times to obtain a mean weight. It is important to note that the accuracy of the check weight is not critical as it is for measurement of drift only. The repeat weighing’s and mean are to be recorded on the Balance Drift Log.

NOTE: *A mass mid-range for the balance in the question would be suitable.

      - The range of the balance is provided on the annual calibration certificate. Suggested balance weights for the balances in CMC Analytical Services are detailed above.

    - Weighing by Tare

      - Ensure the balance is clean, level and the doors are closed.

      - Perform the internal calibration.

      - Place an empty weighting container on the balance pan, press the tare button.

      - Transfer the substance to be weighed to the container, record/print the reading when stable.

    - Weighing by Difference

      - Ensure the balance is clean, level and the doors are closed.

      - Perform the internal calibration.

      - Place an empty weighing container on the balance pan, record/print the stable reading (w0).

      - Transfer the substance to be weighed to the container, record/print the stable reading (w1).

      - Transfer the weighed substance to the required container (e.g., volumetric flask).

      - Replace the weighing container on the balance pan, record/print the stable reading (w2).

      - If the substance was transferred out of the weighing container quantitatively, by washing, the weight of the substance is obtained by (w1-w0).

      - If the substance was not transferred out of the weighing container quantitatively, the weight of substance is obtained by (w1-w2).

      - Clean the balance as described below, after use and between samples and standards.

  - Maintenance of Balances

    - Cleaning

      - Remove the pan and any spill tray, according to the manufacturer’s instructions.

      - Clean the pan, tray and interior of the weighing compartment, then, if necessary, the exterior of the balance and surrounding work areas.

      - Use a tissue soaked in water/detergent mix or a suitable cleaning agent to remove contamination. Dispose of the tissue safely.

      - Thoroughly dry the balance.

      - Check the balance is still level and that reading is stable.

      - Complete the logbook as appropriate.