Calibration and Preventive Maintenance Program

Table of Contents

- Introduction	2

- Definitions/Acronyms	2

- References/Associated Documents	3

- Responsibilities	4

- Calibration and Preventive Maintenance Program	5

- Calibration and Preventive Maintenance Schedule	7

- Decommissioning of Equipment	9

- Labeling of Equipment	10

- Records	10

- Summary of Changes	10

- Introduction

  - Purpose

The purpose of this document is to establish an overall program for purchasing, identifying, maintaining, monitoring, calibrating and qualifying equipment used by LGC Bio search Technologies.

  - Scope

This procedure applies to LGC Bio search Technologies equipment used for the development, production, evaluation and storage of LGC Bio search Technologies products. This includes, but is not limited to, new, relocated, changed/modified, non- conforming or decommissioned equipment.

- Definitions/Acronyms

  - Asset Control Number – A unique identification number that facilitates tracking of calibration or maintenance schedules and other equipment documentation.

  - Calibration – The comparison of a measurement standard or instrument of a known accuracy with another standard or instrument to detect, correlate, report or eliminate by adjustment, any variations in accuracy of the instrument/equipment being evaluated.

  - Change – Any change to equipment. This includes, but is not limited to, equipment which was missed during the regularly scheduled calibration period, equipment that now requires calibration and/or validation/qualification, equipment that has been damaged or compromised subsequent to calibration, equipment requiring operational parameter changes, software changes or upgrades, or decommissioned equipment being reactivated.

  - Decommission – Process through which equipment or instruments are removed from service on a permanent basis. While decommissioned, equipment or instruments may not be used.

  - Equipment – Any machine, electronic or mechanical device used in the development, manufacture, quality control or storage of LGC Bio search Technologies products. This also includes any device used to calibrate or monitor this equipment.

  - Maintenance / Preventive Maintenance – The care and servicing by personnel for the purpose of maintaining equipment and facilities in satisfactory operation condition by providing for systematic inspection, detection, and correction of incipient failures either before they occur or before they develop into major defects.

  - Relocation – Movement of equipment or instruments from one designated area to another. Movement of no-portable equipment, such as balances, analyzers or refrigerators within the same designated area may also be considered “relocation”, since this movement may require recalibration or re-qualification of the instrument or equipment.

  - Retired Status – Permanent removal of equipment from service; disposal.

  - Use Status – Indication of instrument/equipment availability for use. Status may be “In Use” (Available for use), “In Storage” (not available for use), “Discarded” (permanently removed), “Not in Use” (restricted from use), or “Resold” (permanently removed). Status is tracked through the asset control database.

  - System Suitability – Verification that a system is suitable for its intended application.

- References/Associated Documents

  - 21 CFR Part 820, Medical Device; Current Good Manufacturing Practice (cGMP) Final Rule; Quality System Regulation, Food and Drug Administration, Federal Register.

  - ISO 9001, Quality management systems -Requirements, International Organization for Standardization.

  - ISO 13485, Medical Devices – Quality management systems – Requirements for Regulatory Purposes, International Organization for Standardization.

  - ICH Q7 Good Manufacturing Practice Guidance for Active Pharmaceutical Ingredients, Guidance for Industry

  - ICH Q9 Quality Risk Management

  - ICH Q10 Pharmaceutical Quality System

  - FM-202698.02 Equipment Decommissioning Form

  - SOP-202415 Supplier Qualification Procedure

  - SOP-203397 Record Retention Policy

  - WI-201141 Outsourced Preventive Maintenance

  - WI-205661 Controlled Temperature Unit (CTU) Defrost and Preventive Maintenance Procedure

  - WI-201868 Out of Calibration Review

  - WI-202665 Outsourced Calibration Service

  - WI-203094 Commissioning of Manufacturing and Test Equipment

  - WI-203104 Quality Tagout Instructions

  - WI-250139 Equipment Tagout Instructions, Risskov

  - WI-205337 Instructions for Use of the FileMaker Pro Asset Control System

  - SOP-208943 PSOP130 Equipment Qualification and Commissioning Procedure

  - SOP-208891 PSOP032 Validation Master Plan

- Responsibilities

  - Area Management or System Owner

    - Develop new calibration and/or maintenance procedures based on manufacturer recommendations and intended use, as applicable.

  - Process Qualification

    - Responsible for ensuring proper calibration and maintenance programs have been established.

    - Ensure new calibration and/or maintenance procedures have been developed and made effective prior to completion of equipment commissioning.

    - Responsible for reviewing and ensuring that maintenance, calibration or repair was conducted properly.

    - Store and maintain all equipment and instrument records.

    - Schedule outsourced or in-house calibration and maintenance activities, as applicable.

    - Responsible for reviewing and approving maintenance and calibration records.

    - For Germany, site-based Process Qualification designates support 4.2.1 to 4.2.6.

  - Quality Assurance

    - Responsible for auditing Quality System Records for equipment within individual departments.

    - Responsible for reviewing and approving maintenance and calibration records.

    - Ensure calibration and maintenance procedures meet regulatory requirements.

- Calibration and Preventive Maintenance Program

  - All inspection, measuring and testing instruments and/or equipment used in the manufacture or storage of components at LGC Bio search Technologies are considered critical and require the establishment of a calibration and preventive maintenance program dependent upon the equipment manufacturer’s recommendations and intended use.

    - Other programs performed at prescribed intervals may also be required: system suitability testing, temperature verification, system checks, etc.

  - All critical instruments must be selected to meet the accuracy and precision requirements necessary for their intended use. Calibration test data must be evaluated to ensure that these requirements are being maintained.

  - Instruments and equipment must be formally identified with a unique asset control number as a means of traceability. Asset numbers are assigned by Process Qualification or Process Qualification designates for Germany and Denmark.

    - Computerized Equipment Profiles are stored for each item in Asset Control, containing the following data:

      - General information about the item. Serial number, model number, Location (including site, room number and Department).

      - A program of required maintenance items, including calibration, maintenance, verification, system suitability (if applicable), with links to the appropriate Work Instructions for each procedure.

      - A maintenance log. Maintenance logs contain any calibration, maintenance, verification, system suitability, tag outs, service reports and applicable qualification performed on the instrument/equipment.

  - Area Management or System Owner shall work with Process Qualification to determine if the instrument requires calibration and, if so, the calibration frequency.

  - Area Management or System Owner shall work with Process Qualification to determine if the instrument requires Preventive Maintenance, as established by the instrument manufacturer. If Preventive Maintenance is required, Process Qualification will work with Instrumentation to execute a service contract and manage the Preventive Maintenance schedule.

  - Area Management or System Owner shall work with Process Qualification to determine if an instrument requires monitoring such as storage equipment or other equipment requiring monitoring of condition or use.

  - Cycle times for calibration, maintenance, verification, system suitability or other activities shall be stated in the Program tab of Asset Control.

  - Calibration and Maintenance Programs for general lab equipment shall be established according to the parameters below unless otherwise stated in respective asset documentation:

| Type of Equipment | Frequency of Calibration or Maintenance |
| --- | --- |
| Balances and scales measuring less than 50kg | 6 months |
| Balances and scales measuring more than 50kg | Annual |
| Balances, CMC and Therapeutics | Annual |
| Column Heaters | Annual |
| Spectrophotometers | Annual |
| Temperature Sensors | Annual |
| Weight Sets | Annual |
| Hoods: BSC, Laminar Flow or Fume | Annual |

| Pipettes (by Location) | Frequency of Calibration |
| --- | --- |
| Petaluma-RUO | 6 months |
| Petaluma-Therapeutics | 6 months |
| Petaluma-R&D | Annual |
| Novato | 6 months |
| Risskov-ISO9001 | Annual |
| Risskov-GMP | 6 months |
| Steinach | Annual |

  - Calibration and/or Preventive Maintenance shall be performed by qualified in-house personnel or approved service providers.

  - Calibration must be performed using equipment certified by or traceable to nationally recognized standards, such as NIST. Where standards do not exist, the equipment used to calibrate must be thoroughly documented. Instruments and standards used for calibration must have a higher degree of accuracy than the instrument or equipment being calibrated.

  - Written procedures, which identify the measurements to be made, the required accuracy, the appropriate inspection and the measuring and testing equipment, are required.

    - Associated forms detailing the actions performed are submitted to Process Qualification for review and uploading to Asset Control.

    - Maintenance with no associated form or approval requirements may be logged by the individual performing the maintenance.

  - Stickers detailing the calibration, maintenance, system suitability, verification or other activity are issued by Process Qualification after the associated documents have been reviewed for completion and approved.

    - FM-203104.01 Equipment/Instrument Status: Attention to Operator may be attached to the equipment/instrument after the required activity has been performed and submitted to Process Qualification for approval and issuing of sticker.

  - For Germany, stickers are issued by the site lead after completion and approval.

    - FM-203104.07 Attention to Operator STCH may be attached to the equipment/instrument after the required activity has been performed and submitted to site lead for approval and issuing of sticker.

  - Area Management or System Owner will be notified when accuracy and precision requirements are not met or when the calibration or PM due date has been exceeded, as testing may be adversely affected. The instruments/equipment found to be out of tolerance, or which have exceeded the next calibration or PM due date, shall be tagged out according to WI-203104, Quality Tagout Instructions (WI-250139 for Denmark) if they have exceeded the next due date by longer than the established parameters.

    - Equipment found to be out of calibration or specification shall be logged according to WI-201868 Out of Calibration Review.

- Calibration and Preventive Maintenance Schedule

  - The Calibration and Preventive Maintenance Schedule shall be monitored by Process Qualification, or a Process Qualification designate (Germany, Denmark).

    - Scheduling of outsourced calibration, maintenance, and other activities established in the Maintenance Record of Asset Control will be performed by Process Qualification, Instrumentation, Facilities, and a Process Qualification designate (German, Denmark) as applicable.

|  | It is the responsibility of all laboratory personnel to use equipment current on calibration, maintenance or other required activity. |
| --- | --- |

  - Laboratory personnel shall visually check stickers attached to the instrument/equipment and verify they are current. If the sticker is not current, laboratory personnel shall then verify in Asset Control.

    - Laboratory personnel shall notify their Area Manager or Supervisor of equipment that is out of calibration, maintenance, or other required activity.

  - Each month Process Qualification or designate shall review the database for upcoming calibration and PM due dates. Reminders will be sent to the Area Manager or Supervisor.

  - Equipment requiring calibration shall be calibrated by the calibration due date. Any equipment out of specification “as found” will be investigated prior to returning to service; specifically, the impact of the equipment operating out of calibration will be assessed for all materials and products produced or stored during that time. This investigation is conducted with the DNR process.

  - System Suitability Testing shall be performed by the SST due date.

|  | Any equipment or instrumentation that is beyond its calibration or SST due date must be tagged out from use in accordance with WI-203104, Quality Tagout Instructions (WI-250139 for Denmark). This action prevents the use of equipment or instrumentation in the production of products. Any product manufactured with equipment or instrumentation outside of its calibration or SST due date is considered nonconforming and a nonconformance report must be initiated. |
| --- | --- |

  - Equipment requiring Preventive Maintenance shall be serviced no later than one (1) month from the due date listed in Asset Control.

    - For example, equipment with Preventive Maintenance due date of 01 October 2014 may be serviced no later than 01 November 2014 to remain in compliance with this procedure.

    - This window allows for appropriate scheduling of Preventive Maintenance activities, and instruments/equipment past their due date but within the one (1) month window is considered to be in compliance.

  - Equipment/Instruments requiring a Revalidation Assessment based on the scheduled frequency determined by completion of QL-206673 New Equipment/Instrument Initial Validation and Impact Assessment Record shall be assessed no later than one (1) month from the due date listed in Asset Control.

|  | Any equipment or instrumentation that is beyond its PM or Revalidation Assessment due date plus one (1) month must be tagged out from use in accordance with WI-203104, Quality Tagout Instructions (WI-250139 for Denmark). This action prevents the use of equipment or instrumentation in the production of products. Any product manufactured with equipment or instrumentation outside of its PM or Revalidation Assessment due date plus one (1) month is considered nonconforming. |
| --- | --- |

  - Equipment/instruments should continually be observed for proper operation and condition. At any time if equipment/instruments are observed to be malfunctioning, a Facilities or Instrumentation ticket will be opened, and repair or maintenance will be scheduled.

    - Germany and Denmark do not utilize a ticketing system.

  - It is the responsibility of laboratory personnel to observe their CTUs for the existence of frost. Defrosting of CTUs is performed by laboratory personnel or Lab Support, as applicable. See WI-205661 Controlled Temperature Unit Defrost and Preventive Maintenance Procedure for instructions.

    - CTUs shall have no more than 1” of frost buildup at any time.

    - Prior to maintenance, the inventory within storage equipment must be maintained per label requirements and will only be replaced when the equipment equilibrates back to a consistent condition following maintenance. Handling of inventory and equilibration will be documented as part of the maintenance record.

  - Equipment requiring significant, unscheduled maintenance will be reviewed for re- validation according to WI-207663, Equipment/Instrument Re-Validation (Re-Qualification) Assessment Record.

  - If utilizing an outsourced service provider, Process Qualification (or designate), Facilities or Instrumentation shall notify the supplier of the equipment documentation requirements.

    - Service providers shall be evaluated according to:

      - SOP-202415, Supplier Qualification Procedure

      - SOP-250032 Supplier Qualification Procedure, Risskov

      - SOP-208367 Lieferanten management LGC Bio search GmbH, Steinach

- Decommissioning of Equipment

  - Any equipment/instrument intended for decommissioning shall undergo a decommissioning review process. Decommissioning lab or research equipment does not include computers and computer-related devices (e.g., printers, keyboards, monitors, workstations, etc.) as these computers and computer related devices shall be managed by IT.

    - FM-202698.01 Equipment Decommissioning Form is required for all equipment/instruments slated for removal. Completed forms are submitted to Process Qualification.

    - Wherever applicable, equipment/instruments to be decommissioned shall be de- energized if possible before long-term storage or disposal. For most devices, unplugging the device from its power supply will suffice. In other cases, removing pressure from a pneumatic system, gas storage system, or removing charge from a capacitor may be required. De-energizing systems prior to long-term storage or disposal is a crucial step in ensuring safety.

    - Wherever applicable, all equipment/instruments with reagents shall have all chemical or radioactive substances removed and placed in appropriate storage or disposal containers. If there are further decontamination procedures or precautions, coordinate with EH&S to begin the decontamination process. EH&S shall be contacted if the substances require special procedures or precautions for storage or disposal. In the case of a refrigerator or freezer, refrigerant shall not be discharged prior to disposal. Appropriate disinfectants may be utilized in cases where biological materials were in use. In most cases, a 70% ethanol in water solution, a 1:10 bleach in water solution (made fresh), or commercially available disinfectant whose use is explicitly stated as is deemed appropriate for this purpose may be used if used as directed. Note, this does not apply to Risskov.

- Labeling Equipment

  - Accurate labels must be attached to the equipment/instruments undergoing calibration, maintenance, system suitability, verification or other activity.

  - The label will contain the following information:

    - Description of calibration, maintenance, system suitability, verification or other activity.

    - Name of operator performing activity.

    - Date performed and date due.

    - Asset ID.

- Records

  - All calibration, maintenance, system suitability, verification or other activity will be reviewed by Process Qualification and QA.

  - Submitted documents shall be reviewed, signed, and scanned to .pdf by Process Qualification (or designate for Denmark) before logging record in Asset Control. If no signature line is provided, Process Qualification shall initial and date in a visible location on the document.

  - If no asset number is listed on the document, Process Qualification shall write the asset number in a visible location with their initials and date.

  - Documentation will be electronically approved, and physical copies will be stored by Process Qualification.

  - Records shall be retained according to SOP-203397, Record Retention Policy.

- Summary of Changes

| Revision | Description of Change | CRR No. |
| --- | --- | --- |
| 16.0 | Updated references to include WI-205661 in section 3.11; added requirement for review of revalidation of equipment in section 5.2.7; added sections 5.5.2 – 5.5.4 expanding requirements for stored records. | 2019-0825 |
| 17.0 | Clarified roles and responsibilities in section 4; clarified guidelines centered on Asset Management and creation of Cal/PM schedule, added verbiage around activities not covered under Calibration/Preventive Maintenance in section 5; clarified responsibilities around ensuring instruments/equipment are current on Calibration/Preventive Maintenance in section 6; removed unnecessary instructions for decommissioning in section 7; clarified label contents in section 8; updated roles and responsibilities around record keeping of Asset Records in section 9; replaced QA/Calibration with Process Qualification throughout document. | 2021-1709 |
| 18.0 | Section 5.8: Updated Petaluma Therapeutics pipette calibration frequency from 3 months to 6 months. | 2021-1925 |